Hotspin 2
Hotspin 2
Implementation and testing of a new high temperature spinner magnetometer
Michael Wack
[Speaker Notes: Part of my diploma thesis. Original concept developed by Juergen Matzka. Main talk about the instrument, but first motivation ...]

---
Motivation
Motivation
Most paleomagnetic data by stepwise thermal demagnetization
Unnoticed effects in stepwise demagnetization due to alteration of magnetic minerals
Classic interpretation: 3 components including one with magnetite
Hotspin shows mineral unblocking at 350°C ⇒ titano-maghemite and P- resp. N-type (Neel) behavior
Heating
Cooling
1600
Data from
M. Alexutin
1200
Hotspin 1
800
400
0
0
200
400
600
0
100
200
300
400
500
600
Temp [°C]
Temp [°C]
All data from ODP sample 9-0079-017R-01w 102-106
M [arb. units]
130°C
220°C
300°C
350°C
400°C
500°C
600°C
M [arb. units]

---
Plannend features
Planned Features
Temperatures up to 600°C
Full 3D vector of magnetization *
Several specimen at once *
Sensitivity approx. 50 mA/m
Rotation around z-Axis
Possibility to apply magnetic field in z-direction for experiments (TRM, palaeointensities) *
Inch cores *
User friendly software, sophisticated analysis functions *
* Improvement over Hotspin 1
[Speaker Notes: Current status heating must be optimized sensor arrangement will change]

---
Principle of Operation
Principles of Operation
Off-axis spinner magnetometer
Specimens and temperature sensor moving on circular (R=15cm) path
Oven is also circular
Fluxgate sensors (2 for each direction in space – differential measurement eliminates errors due to background fields)
External field compensated by Helmholtz coils

---
off-axis vs. normal
Off-axis vs. normal spinner magnetometer
Normal spinner
Off-axis spinner
Expected signals at sensor
Sinusoidal
Similar to those measured over magnetic anomalies in the subsurface
Data interpretation
Phase and amplitude correspond to the direction and the intensity of the magnetic moment.
Functions calculated from the dipole formula (see later) must be fitted to the measured curves. The determined coefficients are the sought-after components of the magnetic moment.

---
Setup
Setup
Built of non-magnetic materials
All electrical powered stuff away from samples and sensors
Heating by hot air (details later)
Fluxgate sensors (geometrical details later)
Low pass filter to reduce 50Hz and tram noise

---
Heatgun
Heat Gun

---
Heating
Heating Ring
Hot air runs through copper pipes on the in- and outside of the ring
Samples are primarily heated through radiation
Good temperature dispersion (copper is a good conductor of heat)
No magnetic effects
No dangerous voltage

---
Heating Curves
Heating Curves
Heat accumulation directly behind the heat gun due to large flow resistance in the ring ⇒ improve air flow ⇒ draw air out of the ring
Test with vacuum cleaner produced very promising results
Vacuum cleaner turbine was modified for hot air
Heating the whole ring produces higher temperatures than heating a half ⇒ power of heat gun is not the problem
Next try will be the vacuum cleaner turbine with 500W, 13500 rpm
18.10.05 heat gun only
27.10.05 heat gun only, better adapter
28.11.05 radial fan 10W
23.12.05 real vacuum cleaner 1000W
not continuously *
09.01.06 vacuum cleaner turbine 10W *
10.01.06 vacuum cleaner turbine 100W
approx. 6500 rpm *
12.01.06 vacuum cleaner turbine 100W
approx. 6500 rpm
400
* half the ring heated
300
200
100
0
0:00
1:00
2:00
Time [h]
Temperature [°C]
[Speaker Notes: As far as I know no other instruments with heat gun]

---
Electronics schematic
Electronics
Simple lowpass filter will be replace by 5th order filter with special purpose ICs (MAX280)
Signal limiter to +-2.5V to protect the Analog Digital Converter (ADC)
Other devices are operated manually (Helmholtz coils, actuation, exhaustion)
Temp.
sensor
Heat gun
Sensors
Eurotherm Controller
Fluxgate
electronics
Fluxgate
electronics
PC
control software
and
data storage
Lowpass filter
10Hz
Signal limiter
Multiplexer
C
8051
ADC
Position Encoder
RS232
RS232

---
Microcontroller
Microcontroller Board
What does it do?
  - Decodes encoder signals to get the sample position
  - Controls ADC (3 differential signals are sampled in alternating timeslices during each 1/500 rotation; see next slide)
  - Output: differential sensor signals and an absoute position via RS232
Compact unit including ADC (differential, 24 bit) and Multiplexer from Silicon Laboratories
8051 Compatible @ 24.5 MHz
RS232 interface for programming and data exchange
Program written in C

---
ADC
ADC Operation
500 steps / rotation
Number of samples per step depends on the rotational speed
Best usage of time available for sampling data
Best temporal average
Noise reduction
X
Y
Z
X
Y
Z
X
Y
Z
1ms

∅X
∅Y
∅Z
output
6 rpm ⇒ 20 ms

---
Mathemetical Basis
Mathematical Basis
*
vec B( vec r) = {3 vec r left ( vec r cdot vec m right ) - vec m r^2} over {r^5} left [ 10^7 "T" right ]
Coordinate system
View in z direction
vec m = "magnetic dipole moment" ~ left [ "A" "m"^2 right ]
vec r = vec s - vec p = left ( matrix{r_x##r_y##r_z} right ) = left ( matrix{s_x - R cdot cos %PHI ## s_y - R cdot sin %PHI ## s_z} right )
"Magnetization" ~~ vec M = {vec m} over V ~left ["A" over "m" right ]
* Reference: Fließbach, Elektrodynamik, 2000
0 x y %PHI R = 15 "cm" vec r "sensor" vec s "specimen" vec p

---
More Mathematics
More Mathematics
vec B( vec s, %PHI) = breve I cdot vec m = 1 over r^5 left( matrix{ 2 r_x^2 - r_y^2-r_z^2 # 3 r_x r_y# 3 r_x r_z## 3 r_y r_x# 2 r_y^2 - r_x^2-r_z^2# 3 r_y r_z## 3 r_z r_x#3 r_z r_y #2 r_z^2 - r_x^2 - r_y^2} right ) cdot left ( stack{m_x#m_y#m_z} right )
Column:	magnetization component
Row:	mounting direction
e.g. IZy Intensity of magnetization in y-direction
measured by a sensor mounted in z-direction
Expected signal from a sensor mounted in x-direction located at s:
vec B_x( vec s, %PHI) = I_Xx m_x + I_Xy m_y + I_Xz m_z = 1 over r^5 left( (2 r_x^2 - r_y^2 - r_z^2 ) m_x + 3 r_x r_y m_y + 3 r_x r_z m_z right )
y and z analog

---
Optimized Arrangement
Optimized Sensor Arrangement
New arrangement is mechanical more easily to buid and delivers better data.
old arrangement
new arrangement

---
Where to place the sensors
Optimized Sensor Positions
Desirable
  - Big signal amplitude
  - Good seperation of magnetization components
  - Small distance between differential sensors, because of inhomogeneous background field variations
Problem: many degrees of freedom (position in space, orientation, combination with second sensor)

---
Optimized Sensor Positions
Optimized Sensor Positions
Simplifications:
  - Sensor orientation parallel to the principal axes
  - Sensors arranged in one plane (orthogonal to the path of specimen)
  - Sensors as close as possible to specimen
For a given sensor position s one can calculate the amplitudes of the Iij().
  - Amplitude( Iij()) = max( Iij()) - min( Iij()) = Aij

---
Amplitude plots
RGB amplitude plots
A RGB-Color representing sensitivity to the different magnetization components can be assigned to each sensor position.
There is no position to measure x or y separately.
y
z
x
e.g.: Color values for sensor mounted in X direction (left plot):
color red {A_Xx( vec s)} ~~ color green {A_Xy( vec s)} ~~ color blue {A_Xz( vec s)}
z [m]
x [m]
0.1
0.1
0.1
0.2
0.2
0.2
0.045
-0.045
X
Y
Z
Sensor mounting direction
Color values are scaled logarithmic because of 1/r3 law

---
sensor geometry
Sensor geometry
signal amplitudes of single sensors
0.10
0.11
0.12
0.13
0.14
0.15
0.16
0.17
0.18
0.19
0.20
0
2000
4000
6000
8000
12000
14000
16000
18000
X,Y: z = -0.045 m
Z:	z = -0.056 m
X x
X y
X z
Y x
Y y
Y z
Z x
Z y
Z z
-0.04
-0.03
-0.02
-0.01
0
0.01
0.02
0.03
0.04
0
2000
4000
6000
8000
10000
12000
10000
-0.045
0.045
x [m]
z [m]
X1
Z1
Y1
y = 0
X:	x = 0.211 m
Y,Z:	x = 0.200 m
A
30 mm
8 mm
11 mm

---
old vs. new
Old vs. New Sensor Arrangement
differential signals of two sensors
x component of
magnetic moment
y component of
magnetic moment
z component of
magnetic moment
Y
Z
X
Mounting direction: X
Mounting direction: Y
Mounting direction: Y
1000
6000
0
5000
8000
-1000
4000
-2000
3000
6000
-3000
2000
I_{ij} left [ {"V" "s"} over "A" right ]
-4000
1000
4000
0
-5000
2000
-1000
-6000
-2000
-7000
0
-3000
-8000
-4000
-9000
-2000
-5000
-10000
-6000
-4000
-40
-30
-20
-10
0
10
20
30
40
-7000
-6000
-40
-30
-20
-10
0
10
20
30
40
Mounting direction: X
-8000
Mounting direction: Y
9000
6000
8000
5000
7000
4000
-40
-30
-20
-10
0
10
20
30
40
6000
3000
5000
2000
Mounting direction: Z
4000
1000
3000
0
5000
2000
-1000
4000
1000
-2000
3000
0
AZy
-3000
2000
-1000
-4000
1000
-2000
-5000
0
-3000
-6000
AYx
-1000
-4000
-7000
-5000
-2000
-8000
-40
-30
-20
-10
0
10
20
30
40
-40
-30
-20
-10
0
10
20
30
40
-40
-30
-20
-10
0
10
20
30
40
"encoder steps" ~~ %PHI
New sensor
arrangement
Old sensor
arrangement

---
Software
Software
Made with Labview
Keep all raw data
Post stacking
Classic plots (Zijderveld, Stereo, Decay)
Automatic data selection to eliminate disturbed signals (tram, machines in the workshop)
Programmable temperature ramps and ranges

---
Look yourself
Look yourself ...
... now in the basement.
Room K01
one door left of the VFTB